From:	Pomije, Brian D. (LCDR)

Sent:	Wednesday, July 05, 2006 1:12 PM

To:	Burke, Gerald; Clark, Dorothy M. (CIV); Ely, Karen D. (CIV); Fairchild Gregory R LT; Feightner, Michael P. LT; Fong, Raynard K   CAPT; Frampton, Scott W  CIV; Giel, Daniel T. (CTR); Goodwill, Michael E CTR NNMC; Gorsuch, Gregory  M. CAPT; Harden, Gary W; Houston, Tony L. HM2; Johnson, Gabrielle M.; LaFontaine, Richard L CT; Loscocco, Michele F CDR NNMC; Minette, Christopher E. (LCDR); Ozeroglu, Muhammed A  LT; Pomije, Brian D. (LCDR); Scott, John H., CIV.; Simays, Andrew E., CDR; Smith, Glenee L. (Civ); Solanki, Haresh P; Tarver, James K. (CDR); Van Way, Rick L  LCDR (Ret); Younes, Georges  E LCDR

Cc:	Pomije, Brian D. (LCDR); Ely, Karen D. (CIV); Clark, Dorothy M. (CIV); Hernandez, Jose E. (CIV); Fallon, Ann P. (CAPT); Earles, Marvin R LCDR ; Glennon, Brendan K. CDR BUMED; Lino CDR Fragoso (E-mail) (Lino.Fragoso@navy.mil); Hutto, Robert CAPT; Miller, David B, CAPT (NMSC); Nazareno, Felix M GS; Netzer, Patricia H. (CAPT); Styles, Donna L (CAPT)

Subject:	NRC Inspection Value Added Finding: Radiopharmacy Air Sampling and Computer Program Problems

Medical Radiation Safety Officers and NRMP Permittees,

The Nuclear Regulatory Commission (NRC) has recently released the attached letter titled: "NRC Inspection Value Added Finding: Radiopharmacy Air Sampling and Computer Program Problems (STP-06-050)".

In summary, an NRC inspector identified during an inspection of a nuclear medicine radiopharmacy, that a licensee was inaccurately calculating radioactive air concentrations in an unrestricted area.  Specifically, due to a systematic error when evaluating results of air sampling measurements for radioiodine (I-131), the value used for the air exchange rate of the air sampler was a factor of 10 off, resulting in values that potentially could have exceeded federal limits.  In addition, the inspector also discovered errors in the licensee's computer program that was used to ensure that radioactive material possession limits as specified in the facility's license, were not exceeded.

The issues addressed in the letter should be discussed by each Command's Radiation Safety Committee (RSC), to ensure similar errors do not occur within Navy medical facilities.

Please contact the Navy Environmental Health Center if you have any questions and/or concerns.

[This e-mail has been sent to all medical radiation safety officers; no response is necessary]

Very respectfully,

LCDR Pomije

Brian D. Pomije, MS, DABR

Lieutenant Commander, Medical Service Corps, USN

Radiation Health Team Leader

Navy Environmental Health Center

620 John Paul Jones Circle (Suite 1100)

Portsmouth, VA 23708-2103

W: (757) 953-0766; DSN: 377-0766

Cell: (757) 651-2814

Fax: (757) 953-0685

This document may contain information covered under the Privacy Act, 5 USC 552(a), and/or the Health Insurance Portability and Accountability Act (pl 104-191) and it’s various implementing regulations and must be protected in accordance with those provisions. Healthcare information is personal and sensitive and must be treated accordingly. If this correspondence contains healthcare information it is being provided to you after appropriate authorization from the patient or under circumstances that don't require patient authorization. You, the recipient, are obligated to maintain it in a safe, secure and confidential manner. Redisclosure without additional patient consent or as permitted by law is prohibited. Unauthorized redisclosure or failure to maintain confidentiality subjects you to application of appropriate sanction. If you have received this correspondence in error, please notify the sender at once and destroy any copies you have made.